## Slide 1: Virtual Linac Validation 2020 TrueBeam 6X FFF

## Slide 2: Measured data from Antti Kulmala (2016)
SSD = 95cm
40 x 40 x 40 cm3 Blue Phantom 2® -water phantom system
Photon PDDs for field sizes: 3 x 3 cm2, 4 x 4 cm2 and 5 x 5 cm2 and CRs for field size 5 x 5 cm2, have been measured with diode detector.
All other photon beam data is measured using ionization chamber.
Estimate that the combination of uncertainty is ≤ 0.5 % (1σ) for all studied photon beams.
Setup uncertainty is estimated to be ± 0.5 mm (for water phantom and detector separately).

## Slide 3: 100 x batch:
Number of threads 30
Number of events 550000
Simulation time ≈ 1h
Full directional biasing used
SSD = 95cm
Field sizes: 3x3, 4x4, 5x5, 6x6, 10x10, 20x20, 40x40
Scoring grid size (x,y,z) = (1mm, 10mm, 1mm)

## Slide 4: Electron beam properties
Electron energy: 6.5MeV
Energy sigma: 0.056 MeV
Spot size: 0.6645mm x 0.7274mm
Angle divergence: 0.0573°
Determined by T. Ikonen

## Slide 5: PDDs

## Slide 20: Dose profiles

## Slide 21
At depth 15 mm

## Slide 22
Field size 5 x 5 cm2 is been measured with diode detector

## Slide 23
At depth 50 mm

## Slide 25
At depth 100 mm

## Slide 27: Convolving the number of voxels corresponding to the detector size (±3mm)

## Slide 28
At depth 15mm without (left) and with convolution(right)

## Slide 31
Without
With convolution +3mm shift
40x40 at depth 50 and 100 mm

## Slide 32: Diagonals 40x40
Depths: 15mm, 50mm, 100mm, 200mm and 300mm

## Slide 33
40x40 at depth 15 and 50 mm
With convolution
Without

## Slide 34
40x40 at depth 100 and 200 mm
With convolution
Without

## Slide 35
40x40 at depth 300 mm
Without
With convolution